# fNIRS GLOSSARY PROJECT

# LIST OF TERMS

# Topic: Hardware

LETTERS A - Z

CHAIR: Samuel Montero-Hernandez (s.monterohdz@gmail.com)

Please read the landing page with instructions first before you move onto editing this document!

LINK: fNIRS_Glossary_LandingPage

| Team Hardware Please leave your name and e-mail address below if you would like to join team hardware |  |
| --- | --- |
| Name | E-Mail |
| Nicholas (Nick) Ross | nross3@nd.edu |
| David Busch | david.busch@utsw.edu |
| Chia-Chuan Yu | chiachuanyu@utexas.edu |
| Edgar Guevara | edgar.guevara@uaslp.mx |
| Talukdar Raian Ferdous | tferdous@cougarnet.uh.edu |
| Franziska Klein | franziska.klein@uol.de |
| Christian Bartkowski | christian.bartkowski@nirx.de |
| M. Atif Yaqub | atif.yaqub@icfo.eu |
| Heidrun Wabnitz | heidrun.wabnitz@ptb.de |
| Musa Talati | musa.talati.20@ucl.ac.uk |
| Turgut Durduran | turgut.durduran@icfo.eu |
| Alessandro Torricelli | alessandro.torricelli@polimi.it |
| Stefan Carp | stefan.carp@mgh.harvard.edu |

### Template (empty copy that can be copied below as needed).

IMPORTANT NOTE: Please maintain this formatting, including the heading style, labels, and any tags used on the terms.

#### [Term] (Format: font 12, Arial, bold)

Definition: (Format: font size 11, Arial, bold) <add definition here> (Format: font size 11, Arial)

Category:
Related terms: 
Abbreviation or Symbol:

Synonym:

Reference(s):

Alternative definition: (if applicable)

Related terms to alternative definition: (if applicable)

Reference(s) for alternative: (if applicable)

Originally drafted by:

Reviewed (or Edited) by: 
Status:

# Terms Table of Content

LIST OF TERMS

Topic: Hardware

Template (empty copy that can be copied below as needed).

[Term] (Format: font 12, Arial, bold)

Terms Table of Content

fNIRS Glossary of Hardware Terms: A - Z

3D-Digitizer

A

Accelerometer

Ambient Light

Avalanche Photodiode

Average Optical Power

B

Broadband NIRS

C

Calibration

Cap

D

Dark Noise

Detector

Detector Optode

Detector Sensitivity

Diffuse optical tomography

Dual slopes

E

F

Fiber optic

Frequency domain multidistance NIRS

Frequency multiplexing

Status: Complete (agreement from at least 5 team members) 

Full width at half maximum of spectral power distribution

Gyroscope

H

Headgear

High Density Source-detector Array

High-Density Diffuse Optical Tomography

I

Imager

Inertial measurement Unit (IMU)

Instrument response function

J

K

L

Light-emitting diode

M

Magnetometer

Measurement Channel

Mobile fNIRS device

Modulation Frequency

Multi Modulation Frequency fNIRS (mFD-fNIRS)

N

Noise Equivalent Power

O

Optical Loss

Optode

Optode array

Optode cover

P

Peak Wavelength

Phantom

Photomultiplier tube

Power

Power supply

Probe

Q

R

Response Time

S

Sampling frequency

Saturation

Sensor

Short-separation channel

Signal Cross-Talk

Silicon Photodiode

Silicon Photomultiplier

Smart Textiles

Source

Source-detector pair

Source-detector separation

Source Optode

T

Time multiplexing

Transmitter

U

V

Vertical Cavity Surface Emitting Laser

W

X

Y

Z

# fNIRS Glossary of Hardware Terms: A - Z

#### 3D-Digitizer

Definition: A three-dimensional (3D) digitizer measures the exact location of specific points on a real-world object and converts this information into a set of 3D points in a coordinate system. It is typically used to record the position of fNIRS optodes on the participant’s head together with reference points or anatomical landmarks. Later this information can be used to obtain MNI coordinates of the optodes or channels by using spatial registration tools.

In contrast to non-contacting 3D Scanners, which utilize technologies such as lasers, sound, or magnetism to scan an entire object or area, 3D digitizers are close-distance systems that use a stylus or articulated arm to mark points on an actual object based on an electromagnetic field.

Category: Hardware

Related terms: spatial registration, 3D Scanner

Abbreviation or Symbol:

Synonym: Digitizer

Reference(s): https://doi.org/10.1016/j.neuroimage.2005.05.019, https://doi.org/10.1109/EMBC.2013.6611270

https://peerj.com/articles/11804/

10.1002/oa.1156

Alternative definition:

Related terms to alternative definition:

Reference(s) for alternative:

Originally drafted by: Vanessa Reindl

Reviewed (or Edited) by:  Talukdar Raian Ferdous, Aykut Eken, Christian Bartkowski, Samuel Montero-Hernandez, Stefan Carp

Status: Complete (agreement from at least 5 team members)

## A

#### Accelerometer

Definition: Accelerometer is a device that measures the acceleration caused by movement and gravity, typically in one or more of the three planes of movement and/or three rotational axes. In fNIRS applications, accelerometers can be used to measure the amount, degree, and timing of head and/or body movements during measurement to regress out corresponding motion artifacts from brain hemodynamics.It can also be used to flag potential artifacts for further analysis and/or, in cases when movement is part of the protocol, to relate the type of movement to the measured signal.

Category: Hardware

Related terms: motion, movement, motion artifact, movement artifact, gyroscope, Inertial Measurement Unit (IMU), sensor, wearable fNIRS, portable fNIRS, fDCS

Abbreviation or Symbol: ACC, ACCL

Synonym: Motion Sensor

Reference(s): https://doi.org/10.1016/j.jneumeth.2015.02.006

https://doi.org/10.1117/1.3606576

https://doi.org/10.1016/j.neuroimage.2019.116472

https://doi.org/10.1016/j.neuroimage.2019.06.021

Alternative definition:

Related terms to alternative definition:

Reference(s) for alternative:

Originally drafted by: Jason Longhurst

Reviewed (or Edited) by:  D. Busch, Edgar Guevara, Aykut Eken, Christian Bartkowski, Samuel Montero-Hernandez, Turgut Durduran, Stefan Carp

Status: Complete (agreement from at least 5 team members)

#### Ambient Light

Definition: Ambient light is any light that does not originate from the emitters of the fNIRS device (e.g., sunlight, neon light). Depending on their nature, some ambient light sources may affect fNIRS measurements. For instance, they may introduce additional noise into the fNIRS signal or fully saturate the detectors.

Category: Hardware

Related terms: Optode Cover

Abbreviation or Symbol: AML (observed in only one manuscript)

Synonym:

Reference(s): https://doi.org/10.1016/j.ijleo.2020.165116

https://doi.org/10.4172/2167-1222.1000258

https://doi.org/10.1016/j.ijleo.2020.165116

Alternative definition:

Related terms to alternative definition:

Reference(s) for alternative:

Originally drafted by: Jason Longhurst, Meryem Yucel

Reviewed (or Edited) by:  Mahmoud Elsherif, Edgar Guevara, David Mehler, Ségolène Guérin, Rickson Mesquita, Michael Lührs, Stefan Carp

Status: Complete (agreement from at least 5 team members)

#### Avalanche Photodiode

Definition: Avalanche photodiodes utilize the avalanche multiplication process to achieve higher electric output. The avalanche photodiode resembles a PIN photodiode except for the insertion of an additional p-type multiplying region. Most of the incident photons are absorbed in the intrinsic (‘I’) region and electron-hole pairs are first generated in this region as in the PIN photodiode. The holes drift toward the p-type region (‘P’) and the electrons toward the n-type (‘N’). The electrons, before reaching the n-type region, pass through a p-type multiplying region. If the electrons have undergone sufficient acceleration on reaching the multiplying region, new electron-hole pairs are generated by the collision- ionization process. The newly generated electron-hole pairs themselves generate more electron-hole pairs. The drawbacks to avalanche photodiodes are the high bias voltage needed, and the temperature compensation necessary for stable operation.

Category: Hardware

Related terms: Photomultiplier tube, Detector, Sensor, Silicon Photomultiplier

Abbreviation or Symbol: APD

Synonym:

Reference(s): E. Hui Pan. Fiber Optics Weekly Update. Information Gatekeepers Inc.

Alternative definition:  An avalanche photodiode (APD) is a type of semiconductor photodetector used to detect low levels of light in applications such as fNIRS. It is based on the principle of avalanche multiplication, which occurs when an electron in the photodiode collides with an atom and generates a cascade of secondary electrons, resulting in a much larger current than would be caused by a single photon.

APDs have higher sensitivity and lower noise than traditional photodiodes. However, APDs also require a high voltage bias to operate, have a bigger frame size, and can be more expensive than other types of photodiodes.

Related terms to alternative definition:

Reference(s) for alternative:

Originally drafted by: OntoNIRS

Reviewed (or Edited) by:  Franziska Klein, Edgar Guevara, Christian Bartkowski, Samuel Montero-Hernandez

Status: Open for review

#### Average Optical Power

Definition: Average optical power is the mean power of an optical signal (e.g. continuous light or repeated light pulses) within a specific period of time.

Category: Hardware

Related terms:

Abbreviation or Symbol:

Synonym:

Reference(s):

Alternative definition:

Related terms to alternative definition:

Reference(s) for alternative:

Originally drafted by: Meryem Ayşe Yücel

Reviewed (or Edited) by:  Heidrun Wabnitz

Status: Pending

## B

#### Broadband NIRS

Definition: Broadband NIRS (bNIRS) is an optical imaging technique that measures in-vivo changes in tissue hemodynamics (oxyhemoglobin and deoxyhemoglobin) and, often, other biomarkers such as  metabolism (e.g., redox state of cytochrome-c-oxidase or (oxCCO)). Tissue light attenuation measurements are performed at a larger number of wavelengths than typical 2-4 wavelength NIRS instruments, e.g., spanning 600 to 1000 nm at 1 nm wavelength resolution, to reduce crosstalk effects and noise.

In contrast to fNIRS systems, bNIRS use broadband light sources, for example, a tungsten halogen light source emitting light between 200 - 2500 nm and a wavelength-dependent detector (e.g., a spectrophotometer). Optical filters may be placed in the source or detector optical path to narrow the spectrum to the region of interest.

Category: Hardware

Related terms: multi-wavelength fNIRS

Abbreviation or Symbol: bNIRS

Synonym:

Reference(s): https://doi.org/10.1117/1.JBO.21.9.091307

https://doi.org/10.3390/metabo12030260

https://doi.org/10.1038/s41598-021-83420-9

https://doi.org/10.1364/BOE.3.000064

https://doi.org/10.1117/12.700894

Alternative definition:

Related terms to alternative definition:

Reference(s) for alternative:

Originally drafted by: Paola Pinti

Reviewed (or Edited) by:  Edgar Guevara, Christian Bartkowski, D Busch, Musa Talati, M. Atif Yaqub, Turgut Durduran, Alessandro Torricelli, Stefan Carp

Status: Complete (agreement from at least 5 team members)

## C

#### Calibration

Definition: The set of operations which establish, under the conditions of the experiment to be run, the relationship between values indicated by a measuring instrument or measuring system, or values, represented by a material measure, and the corresponding known value of a measurand or its derivative. The result of a calibration permits the estimation of errors of indication of the measuring instrument, or the assignment of values to marks on arbitrary scales. The result of a calibration is sometimes expressed as a calibration factor, or as a series of calibration factors in the form of a calibration curve.

Category: Hardware

Related terms:

Abbreviation or Symbol:

Synonym:

Reference(s): G. van der Horn and J. Huijsing. Integrated Smart Sensors: Design and Calibration. The Springer International Series in Engineering and Computer Science. Springer US, 1997. https://link.springer.com/book/10.1007/978-1-4757-2890-3

Alternative definition:

Related terms to alternative definition:

Reference(s) for alternative:

Originally drafted by: OntoNIRS

Reviewed (or Edited) by:  Mohinish Shukla, Heidrun Wabnitz, M. Atif Yaqub, Nick Ross, D Busch, Stefan Carp

Status: Complete (agreement from at least 5 team members)

#### Cap

Definition: Cap is a wearable garment that is put on a participant’s head to stabilize the positions of the optodes (i.e., hold the light sources and detectors in place) and ensure continuous contact of the optodes with the scalp. It affects (long-term) wearing comfort and can be therefore made of various materials, such as stretchable fabrics, rubber, silicone, or foam, and is often designed to fit a specific head size and shape.It typically has built-in holders or holes to insert the probes or modules. A cap also plays an important role in isolating detectors from ambient light.

Category: Hardware

Related terms: headgear

Abbreviation or Symbol:

Synonym:

Reference(s): https://doi.org/10.1364/ao.54.000576

Alternative definition:

Related terms to alternative definition:

Reference(s) for alternative:

Originally drafted by: Ola Dopierala

Reviewed (or Edited) by: Talukdar Raian Ferdous, Franziska Klein, Christian Bartkowski, Edgar Guevara, M. Atif Yaqub

Status: Complete (agreement from at least 5 team members)

## D

#### Dark Noise

Definition: Dark noise refers to the signal detected by a sensor when no input is present. Ideally, a perfect optical sensor should have a noiseless zero output signal in the absence of light, indicating no background signal or noise. In practice, inherent electronic noise, dark current noise, thermal noise combine to produce a non-zero output signal with noise despite a lack of input light. High dark noise levels and/or a large relative variation of dark noise levels can reduce the quality of measurement data.

Category: Hardware

Related terms: correlated noise, after pulsing, dark current noise, dark current, thermal noise, electronic noise, ambient light, background signal

Abbreviation or Symbol:

Synonym:

Reference(s): Radiometric Calibration: Theory and Methods, Chapter  Dark-Noise Analysis, Clair Wyatt

https://doi.org/10.1117/12.790659

Alternative definition:

Related terms to alternative definition:

Reference(s) for alternative:

Originally drafted by: Christian Bartkowski

Reviewed (or Edited) by:  Talukdar Raian Ferdous, Heidrun Wabnitz, D Busch, Nick Ross, Franziska Klein, Turgut Durduran, Alessandro Torricelli, Stefan Carp

Status: Complete (agreement from at least 5 team members)

#### Detector

Definition: Detector is an optoelectronic element that receives light and converts it into electrical signals. In fNIRS, a detector typically refers to a device, such as an avalanche photodiode or a photomultiplier tube that detects the near-infrared light that has passed through the measured tissue.

Category: Hardware

Related terms: Sensor, Source, Photomultiplier Tube, Avalanche Photodiode, Photodiode, Photodetector, Silicon Photomultiplier

Abbreviation or Symbol: Rx, D, DET

Synonym: Receiver, Sensor

Reference(s): National Cancer Institute. National Cancer Institute Thesaurus, 2017. Documentation page available on http://nciterms.nci.nih.gov. https://doi.org/10.1117/12.2037559, https://www.rp-photonics.com/photodetectors.html,

https://doi.org/10.3390/app10186522,

Alternative definition:

Related terms to alternative definition:

Reference(s) for alternative:

Originally drafted by: Christian Bartkowski, Murat Can Mutlu

Reviewed (or Edited) by:  Christian Bartkowski, Murat Can Mutlu, Nick Ross, Edgar Guevara, Heidrun Wabnitz, Stefan Carp

Status: Complete (agreement from at least 5 team members)

#### Detector Optode

Definition: Detector (or detection) optode is a common term which often identifies a part of an fNIRS system that is in contact with the surface of the tissue being measured. A detector optode either directly contains the photodetector (see definition of “detector”) or can hold the distal end of the detector fibers, and may also include a light guide, a lens, an optical filter and a housing.

Category: Hardware

Related terms: Optode, Detector, Motion Sensor, Source, Photomultiplier Tube, Avalanche Photodiode, Photodiode, Photodetector, Silicon Photomultiplier

Abbreviation or Symbol:

Synonym:

Reference(s):

Alternative definition:

Related terms to alternative definition:

Reference(s) for alternative:

Originally drafted by: Meryem Ayşe Yücel, Heidrun Wabnitz

Reviewed (or Edited) by: Nick Ross, Turgut Durduran, Alessandro Torricelli, Samuel Montero-Hernandez

Status: Open for review

#### Detector Sensitivity

Definition: Detector sensitivity refers to the ability of the detector to convert incoming light/photons into an electrical signal. It is defined as the ratio of the change in the output quantity of the detector against the change in the input quantity of the detector (also known as responsivity). It can be formulated as the observed change in the output current of the detector for a change in the power of the incident light in fNIRS devices (unit: Ampere/Watt or Ampere/Lumen).

Detector sensitivity can be affected by several factors, including the detector material, the size and shape of the detector, the quality of the optics, and the electronics used to amplify the detector signal.

Category: Hardware

Related terms: Detector, Detectivity, Responsivity, Noise equivalent power

Abbreviation or Symbol:

Synonym:

Reference(s): https://doi.org/10.1201/9781315217109

Alternative definition:

Related terms to alternative definition:

Reference(s) for alternative:

Originally drafted by: Murat Can Mutlu

Reviewed (or Edited) by:  Franziska Klein, Christian Bartkowski, Usama A Syed, M. Atif Yaqub, Samuel Montero-Hernandez, Heidrun Wabnitz

Status: Complete (agreement from at least 5 team members)

#### Diffuse optical tomography

Definition: Diffuse optical tomography (DOT) refers to the use of a tomographic algorithm to reconstruct, at least (but not limited to)  oxygenated and deoxygenated hemoglobin concentrations volumetrically from multiple spatially distinct measurements. This 3D reconstruction can improve quantification and localization of cortical hemodynamics.

Category: Hardware

Related terms: high-density DOT, diffuse optical imaging (DOI), fNIRS, NIRS

Abbreviation or Symbol: DOT

Synonym:

Reference(s): https://doi.org/10.1117/1.NPh.6.3.035007

https://doi.org/10.1364/AO.45.008142

Alternative definition:

Related terms to alternative definition:

Reference(s) for alternative:

Originally drafted by: Lauren Emberson

Reviewed (or Edited) by:  David Busch, Franziska Klein, Nick Ross, M. Atif Yaqub, Samuel Montero-Hernandez, Turgut Durduran, Meryem Ayse Yucel

Status: Complete (agreement from at least 5 team members)

#### Dual slopes

Definition: Dual slope approach requires at least two sources and two detectors organized in specially structured arrays and either a continuous-wave (CW) or frequency domain (FD) system. The approach gives increased sensitivity to deeper tissue as well as more robustness to experimental errors and variability in opto-mechanical coupling. The spatial gradient of DC amplitude, AC amplitude, or phase vs source-detector distance is referred to as the slope.

Category: Hardware

Related terms:

Abbreviation or Symbol:

Synonym:

Reference(s): https://doi.org/10.1364/JOSAA.36.001743

Alternative definition:

Related terms to alternative definition:

Reference(s) for alternative:

Originally drafted by:  Edgar Guevara, D Busch

Reviewed (or Edited) by:  Nick Ross, Turgut Durduran, Stefan Carp

Status: Open for review

## E

## F

#### Fiber optic

Definition: Fiber optic is a filament of a dielectric material that has transparency to a portion of the optical spectrum and structure so that light is propagated through the fiber. It consists of a central core surrounded by a cladding. In a typical step-index fiber, the core has an index of refraction $$n_{1}$$ that is slightly higher than that of its surrounding cladding $$n_{2}$$. In a ray-optics model, illuminating the end of the fiber starts rays along many possible paths in the core and cladding. A ray whose path meets the core-clad boundary at a small enough angle is reflected back into the core, and thereby, is confined to the core by a series of total internal reflections and such a ray may thus propagate for long distances (>km) in the fiber core. Similar principles apply to graded-index fiber optics for which the transition between $$n_1$$ and $$n_2$$ is gradual.

Category: Hardware

Related terms:

Abbreviation or Symbol: FO

Synonym: optical fiber, glass fiber, optical guide

Reference(s):  Optical Materials: Volume 1. Optical Materials. Taylor & Francis, 1990

https://doi.org/10.1201/9781003210139

Alternative definition:

Related terms to alternative definition:

Reference(s) for alternative:

Originally drafted by: OntoNIRS

Reviewed (or Edited) by:  D Busch, Christian Bartkowski, Nick Ross, Alessandro Torricelli, Samuel Montero-Hernandez

Status: Complete (agreement from at least 5 team members)

#### Frequency domain multidistance NIRS

Definition: By capturing the change in intensity and phase at multiple source detector separations/distances, the frequency-domain (FD) multidistance NIRS technique can estimate absolute values of absorption and scattering of the medium, and subsequently chromophore concentrations. This may involve one or more modulation frequencies. The key here is that these measurements are taken at multiple source detector separations as opposed to one. Ideally, this allows one to obtain the optical properties of the medium under investigation without performing a calibration measurement on a tissue-like medium. Additionally, this allows a simpler processing method requiring linear regression only instead of non-linear optimization to get absorption and scattering of the medium. However, most practical multi-distance frequency-domain multidistance NIRS systems do utilize a phantom calibration of the coupling coefficients between the sources/detectors and media.

Category:  Hardware

Related terms: frequency domain, source detector separations, time domain, continuous wavelength fNIRS

Abbreviation or Symbol: FDMD-NIRS

Synonym:

Reference(s): https://doi.org/10.1117/12.183988

https://dx.doi.org/10.1088/0031-9155/46/1/304

https://doi.org/10.1364/BOE.2.000552

Alternative definition:

Related terms to alternative definition:

Reference(s) for alternative:

Originally drafted by: Nick Ross

Reviewed (or Edited) by:  Franziska Klein, M. Atif Yaqub, D Busch, Turgut Durduran, Heidrun Wabnitz, Stefan Carp

Status: Complete (agreement from at least 5 team members)

#### Frequency multiplexing

Definition: Distinguishing different channels by modulating the light source at non-overlapping frequencies. In fNIRS devices, each light emitter can be codified with a specific frequency modulation through frequency multiplexing, which allows simultaneous detection of all sources at any given detector. The demultiplexing would then be employed to measure the amount of signal detected coming from each source.

Category: Hardware

Related terms: Time multiplexing

Abbreviation or Symbol: N/A

Synonym: NaN

Reference(s): DOI: 10.1117/1.nph.8.1.012101

Alternative definition:

Related terms to alternative definition:

Reference(s) for alternative:

Originally drafted by: OntoNIRS, Rickson Mesquita

Reviewed (or Edited) by:  Franziska Klein, Rickson Mesquita, Guy Perkins, Stefan Carp

#### Status: Complete (agreement from at least 5 team members)

#### Full width at half maximum of spectral power distribution

Definition: Difference of the wavelength between the two points whose corresponding power values are equal and 3 dB lower than the values at each peak wavelength.

Category: Hardware

Related terms: peak wavelength

Abbreviation or Symbol: FWHM

Synonym:

Reference(s): IEC 80601-2-71:2015, 201.3.204

Alternative definition:

Related terms to alternative definition:

Reference(s) for alternative:

Originally drafted by: IEC 80601-2-71

Reviewed (or Edited) by:  Heidrun Wabnitz

Status: No review (ISO/IEC term)

G

#### Gyroscope

Definition: A gyroscope is a device that measures or maintains orientation and angular velocity (in rad/s). It consists of a spinning rotor that detects changes in angular velocity when the frame it is mounted on rotates. Gyroscopes can be used independently or integrated into inertial measurement units (IMUs).

In fNIRS, gyroscopes are often used in combination with an accelerometer mounted at the optode level to track head motion (i.e., position and orientation) during data acquisition.

Category: Hardware

Related terms: Accelerometer, Inertial Measurement Unit (IMU), Motion Artifacts

Abbreviation or Symbol: Gyro

Synonym:

Reference(s): https://doi.org/10.3390/s17102284

Alternative definition:

Related terms to alternative definition:

Reference(s) for alternative:

Originally drafted by: Christian Bartkowski, Franziska Klein

Reviewed (or Edited) by:  Edgar Guevara, Christian Bartkowski, Nick Ross, M. Atif Yaqub, Samuel Montero-Hernandez

Status: Complete (agreement from at least 5 team members)

## H

#### Headgear

Definition: Equipment and materials that hold the sources and detectors in a specific source-detector arrangement with a good scalp coupling at the desired location on the head. Headgear may include optode grids that maintain a fixed source-detector spacing as well as headbands or caps to place the optodes on the head.

Category: Hardware

Related terms: Cap

Abbreviation or Symbol:

Synonym:

Reference(s): https://doi.org/10.1016/j.neubiorev.2009.07.008

https://doi.org/10.1016/j.dcn.2018.01.009

http://dx.doi.org/10.1016/j.dcn.2016.11.005

http://dx.doi.org/10.3389/fpsyg.2011.00222

Alternative definition:

Related terms to alternative definition:

Reference(s) for alternative:

Originally drafted by: M. Atif Yaqub

Reviewed (or Edited) by:  Mahmoud Elsherif; Gregoire Vergotte, Vanessa Reindl, Christian Bartkowski, Simone Cutini, Jason Longhurst

Status: Complete (agreement from at least 5 team members)

#### High Density Source-detector Array

Definition: Arrangement of sources and detectors in an interlaced lattice with the first nearest neighbor distance of up to 15 mm in adults for most implementations. It would be different in an infant head. The term “high density” allows it to be differentiated from the “traditional” sparse optode placement with several cm (~3) between optodes.

Category: Hardware

Related terms: High Density Diffuse Optical Tomography (HD-DOT), Ultra-high density (UHD)

Abbreviation or Symbol: HD

Synonym:

Reference(s): https://doi.org/10.1063/1.5086809

https://doi.org/10.1109/OJEMB.2023.3248492

https://doi.org/10.1117/1.NPh.10.2.025007

https://fnirs.org/2022/01/01/sfnirs2022_34/

Alternative definition:

Related terms to alternative definition:

Reference(s) for alternative:

Originally drafted by: Edgar Guevara

Reviewed (or Edited) by:  Edgar Guevara, Christian Bartkowski, Nick Ross, M. Atif Yaqub, Samuel Montero-Hernandez, Turgut Durduran, Stefan Carp

Status: Complete (agreement from at least 5 team members)

#### High-Density Diffuse Optical Tomography

Definition: High-Density Diffuse Optical Tomography (HD-DOT) imaging systems use dense, as in “High Density Source-detector Array“, often regularly spaced arrays of near-infrared sources and detectors to obtain overlapping measurements of the underlying hemodynamic activity. Compared to conventional multichannel fNIRS systems, the measurements obtained from source-detector pairs across various increments of neighboring distances facilitate higher spatial resolution and depth profiling. This allows a three-dimensional tomographic spatial reconstruction of cortical activity patterns comparable to those obtained using fMRI.

Category: Hardware

Related terms: Diffuse Optical Tomography (DOT), Ultra-high density (UHD), High Density (Source-detector Array)

Abbreviation or Symbol: HD-DOT, HDDOT

Synonym:

Reference(s): https://doi.org/10.1117/1.NPh.2.2.025005

https://doi.org/10.1016/j.neuroimage.2009.07.033

https://doi.org/10.1016/j.neuroimage.2012.01.124

https://doi.org/10.1016/j.neuroimage.2015.05.058

https://doi.org/10.1080/23273798.2017.1290810

https://doi.org/10.7554/eLife.75323

https://doi.org/10.1063/1.5086809

https://doi.org/10.1117/1.3368999

Alternative definition:

Related terms to alternative definition:

Reference(s) for alternative:

Originally drafted by: Aahana Bajracharya

Reviewed (or Edited) by:  Franziska Klein, Christian Bartkowski, Nick Ross, Samuel Montero-Hernandez, Turgut Durduran, Stefan Carp

Status: Complete (agreement from at least 5 team members)

## I

#### Imager

Definition:  In the field of neuroimaging, an imager is a device records and visualizes (live) the spatial distribution of brain activity status (hemodynamic changes). Visualization typically occurs through acquisition software, presenting data in the form of diagrams, 3D models, or topographic maps to monitor the brain activity. They can be categorized as either mobile/wearable imagers or stationary/laboratory devices, depending on their portability.

Category: Hardware

Related terms: Mobile fNIRS device

Abbreviation or Symbol:

Synonym: imager system, imaging system, imaging device, fNIRS imager

Reference(s): https://doi.org/10.1109/NEBC.2003.1215988

https://doi.org/10.1063/1.4704456

https://doi.org/10.1117/12.341439

https://doi.org/10.1007/978-3-319-91287-5_22

https://doi.org/10.1016/j.neuroimage.2012.03.049

Alternative definition:

Related terms to alternative definition:

Reference(s) for alternative:

Originally drafted by: Christian Bartkowski

Reviewed (or Edited) by:  Edgar Guevara, Nick Ross, Franziska Klein, Heidrun Wabnitz, Stefan Carp

Status: Complete (agreement from at least 5 team members)

#### Inertial measurement Unit (IMU)

Definition: An Inertial Measurement Unit (IMU) is a device that measures acceleration, angular velocity, and sometimes the body's orientation, using a combination of accelerometers, gyroscopes, and/or magnetometers. It is often used for motion tracking in robotics, virtual reality, and other applications.

In fNIRS research, IMU can be attached to the subject's head to track head motion during an fNIRS experiment, and this information can be used to remove the effects of head motion on the fNIRS data.

Category: Hardware

Related terms: Motion Correction, Accelerometer, Gyroscope, Magnetometer

Abbreviation or Symbol: IMU

Synonym:

Reference(s): https://doi.org/10.1109/TNSRE.2020.2987888

https://doi.org/10.1186/s12938-018-0554-9

Alternative definition:

Related terms to alternative definition:

Reference(s) for alternative:

Originally drafted by: Franziska Klein, Talukdar Raian Ferdous

Reviewed (or Edited) by:  Talukdar Raian Ferdous, Edgar Guevara, Christian Bartkowski, M. Atif Yaqub, Samuel Montero-Hernandez

Status: Complete (agreement from at least 5 team members)

#### Instrument response function

Definition: A gauge of how the components of a time-domain (TD) fNIRS system affect the experimental data. The IRF is measured by recording  a histogram of photon arrival times using the TD fNIRS system and optodes, with a carefully-designed mechanism replacing the tissue in the measurement optical light path which provides both optical attenuation and sufficient scattering, without broadening the laser pulse itself. The IRF is influenced by the laser pulse shape, the temporal response of the detector and electronics, and pulse broadening due to temporal dispersion in the optics.

Category: Hardware

Related terms: TD fNIRS

Abbreviation or Symbol: IRF

Synonym:

Reference(s): http://dx.doi.org/10.1016/j.neuroimage.2013.05.106

Alternative definition: The IRF characterizes the time resolution of a time-domain (TD) fNIRS system. The IRF is determined by inserting a reference sample between source and detection components of the instrument which replaces the tissue and introduces negligible additional temporal dispersion. The IRF depends on the laser pulse shape, the temporal response of the detector and electronics, and pulse broadening due to temporal dispersion in the optics, and is represented as a convolution of these individual effects.

Related terms to alternative definition:

Reference(s) for alternative: http://dx.doi.org/10.1117/1.JBO.19.8.086010

Originally drafted by: Alessandro Torricelli, Heidrun Wabnitz (alternative definition)

Reviewed (or Edited) by:  D. Busch, Heidrun Wabnitz, Nick Ross, Franziska Klein, Turgut Durduran, Alesandro Torricelli, Stefan Carp

Status: Complete (agreement from at least 5 team members)

## J

## K

## L

#### Light-emitting diode

Definition: Light-emitting diode is a p-n junction, a semiconductor-based optoelectronic device, driven by a current $$I$$ into forward bias. Electrons and holes recombine close to the junction between the p- and n-materials and give out radiation with a frequency determined by the photon energy which in turn is determined by the material’s impurities or band gap. Some fraction of the forward current $$I$$ is then turned into useful light $$L$$ (power) formed from photons with a mean energy $$hf_{m}$$. The quantum efficiency is $$\eta = (eL/hf_{m}I)$$. The more efficient light emitting diodes have a well-defined recombination region (volume $$\Phi$$).

Category: Hardware

Related terms: emitter, source, laser, transmitter

Abbreviation or Symbol: LED

Synonym:

Reference(s): J.E. Carroll. Rate Equations in Semiconductor Electronics. Cambridge University Press, 1990.

Alternative definition:  An LED is a semiconductor device in which light is generated when an electric current is passed through it. The generated light has a wider spectral width and lower output power compared to a laser. However, LEDs have the advantage of being fairly inexpensive and portable.

Related terms to alternative definition:

Reference(s) for alternative:  https://doi.org/10.1111/php.13219

Originally drafted by: OntoNIRS

Reviewed (or Edited) by:  Edgar Guevara, Franziska Klein, Heidrun Wabnitz, Nick Ross, Samuel Montero-Hernandez

Status: Complete (agreement from at least 5 team members)

## M

#### Magnetometer

Definition: A magnetometer is a device that measures a surrounding magnetic field (e.g., the Earth’s magnetic field). Magnetometers are often included in inertial measurement units where they are used as the heading reference and to define the orientation of an object with respect to the Earth. Pairing a magnetometer with an accelerometer and gyroscope can improve the identification and removal of motion artifacts. Magnetic field strength is measured in micro Tesla ($$\mu$$T).

Category: Hardware

Related terms: inertial measurement unit, accelerometer, gyroscope

Abbreviation or Symbol: MAG

Synonym:

Reference(s): https://doi.org/10.1186/s12938-018-0554-9

Alternative definition:

Related terms to alternative definition:

Reference(s) for alternative:

Originally drafted by: Franziska Klein

Reviewed (or Edited) by:  Usama A Syed, Nick Ross, Stefan Carp, Samuel Montero-Hernandez

Status: Open for review

#### Measurement Channel

Definition: A channel in fNIRS is typically the combination of a source and a detector that together produce a single stream of measurement data from a circumscribed portion of underlying brain (or other) tissue. A channel's position is marked mid-way between its source and detector pair in a 3D space, which also indicates the surface projection of the underlying brain area where emitted photons have the highest probability of interacting with brain tissue.

Category: Hardware

Related terms: Short-separation channel

Abbreviation or Symbol:

Synonym: Channel, Source-detector pair

Reference(s):https://doi.org/10.1364/BOE.7.003491

Alternative definition:

Related terms to alternative definition:

Reference(s) for alternative:

Originally drafted by: Mohinish Shukla

Reviewed (or Edited) by:  Edgar Guevara, Heidrun Wabnitz, Mohinish Shukla, Nick Ross, M. Atif Yaqub, Turgut Durduran, Alessandro Torricelli, Meryem Ayşe Yücel

Status: Complete (agreement from at least 5 team members)

#### Mobile fNIRS device

Definition: Mobile devices are lightweight, portable and wearable. Due to their compact and robust design, mobile fNIRS devices can be used outside the laboratory/controlled environment and in everyday life situations, enabling freely moving participants in unrestrained settings.

Category: Hardware

Related terms: Naturalistic environment, Unrestrained settings, Movement

Abbreviation or Symbol:

Synonym: Portable fNIRS, Wearable fNIRS, Mobile Device

Reference(s): https://doi.org/10.1111/nyas.13948, https://doi.org/10.1364/BRAIN.2020.BM3C.2

Alternative definition:

Related terms to alternative definition:

Reference(s) for alternative:

Originally drafted by: Oliver Seidel-Marzi

Reviewed (or Edited) by:  Edgar Guevara, Christian Bartkowski, Usama A Syed, M. Atif Yaqub, Samuel Montero-Hernandez

Status: Complete (agreement from at least 5 team members)

#### Modulation Frequency

Definition: Modulation frequency is the frequency (f) at which the amplitude of light is changed or modulated, in a source in frequency domain hardware.

The light detected in frequency domain hardware has three components: the DC intensity, which is analogous to continuous wave measurement, the AC intensity (I) , which is the intensity of light oscillating at the given modulation frequency, and the phase (ϕ), which represents the phase difference of the emitted and detected modulated light. The phase of the detected light therefore provides information regarding how much it has been delayed (on average) as a result of photon interactions with a medium, i.e scattering. The phase is directly related to the time delay of light, τ , such that ϕ is in the order of ωτ , where ω=2πf and ω is the angular modulation frequency. For a good signal to noise ratio, it is required that ωτ≈1(radians). If ωτ is ≪ 1, there would not be a measurable phase change. If ωτ is ≫ 1, then the AC amplitude decreases and could be below the level of noise.

In the context of imaging biological tissue, time delays are on the order of nanoseconds (for source-detector separations of 2-3 centimeters), which yields optimum modulation frequencies in the order of 100 MHz.

Category: Hardware

Related terms: Frequency domain, Phase

Abbreviation or Symbol: lower case omega

Synonym: Angular modulation frequency

Reference(s):  https://doi.org/10.1364/BOE.467614

https://doi.org/10.1117/3.2219603

https://doi.org/10.1017/cbo9781139029797

Alternative definition:

Related terms to alternative definition:

Reference(s) for alternative:

Originally drafted by: Guy Perkins, Heidrun Wabnitz

Reviewed (or Edited) by:  M. Atif Yaqub, Nick Ross, Stefan Carp

Status: Open for review

#### Multi Modulation Frequency fNIRS (mFD-fNIRS)

Definition: Frequency Domain measurements are typically taken at a single modulation frequency. The three measured signals from these measurements are the direct current amplitude, alternating current amplitude, and phase. These signals have a modulation frequency dependence. Multi Modulation Frequency fNIRS (mFD-fNIRS) combines multiple measurements of FD-fNIRS at different modulation frequencies, either simultaneously or by repeated measurements, each at a discrete modulation frequency. The measurements at discrete modulation frequencies are then used simultaneously for spectroscopic or tomographic reconstruction of optical properties. In current literature, the frequencies typically range from tens of MHz to 1000 MHz and have been shown to increase localization and better image resolution compared to single modulation frequency FD-fNIRS.

Category: Hardware

Related terms: Frequency Domain, Modulation Frequency

Abbreviation or Symbol: mFD-fNIRS

Synonym:

Reference(s): https://doi.org/10.1364/BOE.435913

https://doi.org/10.1364/BOE.467614

https://doi.org/10.1117/1.NPh.8.4.045002

Alternative definition:

Related terms to alternative definition:

Reference(s) for alternative:

Originally drafted by: Guy Perkins

Reviewed (or Edited) by:  Franziska Klein, Usama A Syed, Nick Ross, Turgut Durduran, Stefan Carp

Status: Complete (agreement from at least 5 team members)

## N

#### Noise Equivalent Power

Definition: Noise Equivalent Power (NEP) quantifies the sensitivity of a detector (also often used for the instrument to define the sensitivity of the entire system). It is the minimum input power required to produce a usable output signal in the presence of noise, or in other words, a measure of the weakest optical signal that can be detected. Therefore, it is desirable to have a NEP as low as possible since a low NEP value corresponds to a lower noise floor and, consequently, a more sensitive detector. A low NEP is beneficial even at higher input intensities since it will lead to lower noise characteristics in the output signal.

Technically, NEP is the input power that produces a signal-to-noise ratio (SNR) of one in a 1 Hertz (Hz) bandwidth. For photodetectors it is given in Watts per square root of Hertz (W/√Hz). However, it is also not uncommon to list a system’s sensitivity in terms of its total individual bandwidth in Watts only. Typical dimensions in the field of fNIRS are in the magnitude of pW

To calculate Noise Equivalent Power, it is necessary to know the spectral sensitivity of the photodetector, dark noise performance, measurement bandwidth, and the current-to-voltage conversion factor.

Category: Hardware

Related terms: detector, performance, sensitivity

Abbreviation or Symbol: NEP

Synonym:

Reference(s): 10.3390/s20102831, 10.3390/s21186106, 10.1016/j.neuroimage.2022.119216, Corpus ID: 7011179, https://www.thorlabs.com/images/TabImages/Noise_Equivalent_Power_White_Paper.pdf

Alternative definition:  Noise equivalent power is a metric used to evaluate the sensitivity of a detector, which is the minimum input power required to produce a usable output signal in the presence of noise. The lower the NEP value, the more sensitive the detector is, since a low NEP corresponds to a lower noise floor. NEP is calculated as the input power that produces a signal-to-noise ratio (SNR) of one in a 1 Hertz (Hz) bandwidth, typically measured in Watts per square root of Hertz (W/√Hz). In the field of fNIRS, NEP values are typically in the magnitude of picoWatts (pW).

Related terms to alternative definition:

Reference(s) for alternative:

Originally drafted by: Christian Bartkowski, Edgar Guevara (alternative definition)

Reviewed (or Edited) by:  Edgar Guevara, Murat Can Mutlu, Nick Ross, Stefan Carp, Samuel Montero-Hernandez

Status: Complete (agreement from at least 5 team members)

## O

#### Optical Loss

Definition: Optical loss refers to the reduction in light intensity or power as it propagates through a medium (e.g. tissue or phantom). This loss occurs due to various factors such as absorption, scattering, and reflection of light. Optical loss can be quantified as the ratio between the input optical power (light emitted by the sources) and the output optical power (light received by the detectors after traversing the medium). This ratio, often expressed in decibels (dB), provides a measure of how much the emitted light signal has attenuated or weakened by the time it reaches the detectors.

Category: Hardware

Related terms: functional NIRS phantom

Abbreviation or Symbol: NaN

Synonym:

Reference(s):

Alternative definition:

Related terms to alternative definition:

Reference(s) for alternative:

Originally drafted by: Meryem Ayşe Yücel

Reviewed (or Edited) by:  Heidrun Wabnitz

Status: Open for review

#### Optode

Definition: An optode is the distal end of an fNIRS system which interfaces with the surface of the tissue being measured. Optodes can contain a single source (source optode), a single detector (detector optode), or a source accompanied by a short-separation detector.

Category: Hardware

Related terms: source; detector; emitter; emission, receiver, transmitter

Abbreviation or Symbol:

Synonym:

Reference(s): https://doi.org/10.1016/j.neuroimage.2005.08.008

https://biomedical.njit.edu/sites/biomedical/files/fNIRS%20Optode%20Helmet%20Construct.pdf

Alternative definition:

Related terms to alternative definition:

Reference(s) for alternative:

Originally drafted by: David Busch, Christian Bartkowski, Meryem Ayşe Yücel

Reviewed (or Edited) by:  D Busch, Talukdar Raian Ferdous, Aykut Eken, Edgar Guevara, Christian Bartkowski, Mohinish Shukla, Heidrun Wabnitz, Alessandro Torricelli, Stefan Carp, Alessandro Torricelli

Status: Complete (agreement from at least 5 team members)

#### Optode array

Definition: Relative arrangement of linked sources and detectors on a region of interest.

Category: Hardware

Related terms: Probe design; Optode placement, topographic layout, Montage

Abbreviation or Symbol:

Synonym: Probe array; fNIRS array; Channel configuration; fNIRS-sensor configuration; Source–detector array, fNIRS Montage

Reference(s): https://doi.org/10.1117/1.NPh.5.3.035010

https://doi.org/10.1117/1.NPh.8.1.012101

Alternative definition:  A schematic arrangement of sources and detectors to create a pattern of measurement channels.

Related terms to alternative definition:

Reference(s) for alternative:  https://doi.org/10.1117/1.NPh.8.1.012101

Originally drafted by: Ségolène Guérin

Reviewed (or Edited) by:  Christian Bartkowski, M. Atif Yaqub, Samuel Montero-Hernandez, Nick Ross, Franziska Klein

Status: Complete (agreement from at least 5 team members)

#### Optode cover

Definition: An optode cover is a protective layer of material placed over an optode to prevent ambient light from being recorded by the detectors.

Category: Hardware

Related terms:

Abbreviation or Symbol:

Synonym: light shield

Reference(s): https://doi.org/10.1289%2FEHP2049

Alternative definition:

Related terms to alternative definition:  Ambient Light

Reference(s) for alternative:

Originally drafted by: Aykut Eken, Edgar Guevara

Reviewed (or Edited) by:  Aykut Eken, Edgar Guevara, Christian Bartkowski, Samuel Montero-Hernandez, Nick Ross, Franziska Klein

Status: Complete (agreement from at least 5 team members)

## P

#### Peak Wavelength

Definition: Wavelength where the power is the largest in the spectral power distribution for each of the distinct nominal wavelengths in the light radiated from the emitter optode.

Category: Hardware

Related terms: Full width at half maximum of spectral power distribution

Abbreviation or Symbol:

Synonym:

Reference(s): IEC 80601-2-71:2015, 201.3.214

Alternative definition:

Related terms to alternative definition:

Reference(s) for alternative:

Originally drafted by: IEC 80601-2-71:2015, 201.3.214

Reviewed (or Edited) by: Heidrun Wabnitz

Status: No review (ISO/IEC term)

#### Phantom

Definition: Phantoms, whether material or digital (simulations), serve as models to validate the quantitative results of imaging systems. They are specifically designed to mimic certain in vivo characteristics and the data collected using phantoms can be used as reference values for in vivo measurements.

In the field of fNIRS, researchers use various types of material phantoms:

(1) Solid phantoms which are made of a solid material like epoxy resin or polyoxymethylene containing scattering constituents, such as titanium oxide (TiO2), and absorbing constituents, such as nigrosin. Solid phantoms are commonly static and cannot model dynamic changes within the brain unless an inhomogeneity is being mechanically moved (or thermoelectrically activated) within the phantom.

(2) Semi-solid phantoms which are made of a gel-like material containing scattering and absorbing particles and help test the effects of tissue movement and deformation on fNIRS signals.

(3) Liquid phantoms that consist of a solvent (e.g., water), a scattering agent (e.g., fat emulsion, latex spheres), and an absorber (e.g., ink,  hemoglobin). Glycerol may be added to change the viscosity, simulating a change in blood flow for DCS and similar techniques. They can model dynamic changes, however, these phantoms are typically not stable over long periods (weeks) and layered models (e.g., to simulate the brain and superficial layers) are challenging.

(4) Multi-layer phantoms are made of multiple layers of different materials with different optical properties to simulate the layered structure of human tissue, e.g., solid compartments modeling the superficial layers of the head and a liquid section mimicking the brain layer.

Digital (simulation) phantoms  refer to different means of simulating (often with added   noise) the effects of layers, heterogeneities, optical properties, boundaries and others on the detected signals. They can be used to test different algorithms of signal analysis,  to design new hardware and to aid the interpretation of in vivo data.

Category: Hardware

Related terms:

Abbreviation or Symbol:

Synonym:

Reference(s): https://doi.org/10.1002/nbm.4421

https://doi.org/10.1186/s40658-020-00320-z;

https://doi.org/10.1364/BOE.9.002068;

https://doi.org/10.1088/0031-9155/40/5/016;

https://doi.org/10.1109/TIM.2020.2990261;

https://doi.org/10.1117/1.JBO.20.12.121304

Alternative definition:

Related terms to alternative definition:

Reference(s) for alternative:

Originally drafted by: Anna Exner, Christian Bartkowski, Heidrun Wabnitz

Reviewed (or Edited) by:  David Busch, Usama A Syed, Nick Ross, Franziska Klein, Samuel Montero-Hernandez, Turgut Durduran, Alessandro Torricelli

Status: Complete (agreement from at least 5 team members)

#### Photomultiplier tube

Definition: Photomultiplier tube (PMT) is a photodetection device equipped with a photocathode and electron multiplier utilized for its high gain and low light sensing capabilities. When a photon strikes the photocathode, a single photoelectron is emitted from this site. The photoelectron is then accelerated using electric fields to a structure known as a dynode. After striking the dynode, additional electrons are released which are then sent to another dynode. This process will continue for as many dynodes as contained in the PMT. The resulting photoelectron gain is ~1x106 to 1x107. In analog mode, the output of the device is a current that can be measured directly from the device. Additionally, a PMT can operate in photon counting mode, which, when combined with a pulse height discriminator and timing and counting electronics, may be utilized for TD-fNIRS.

Category: Hardware

Related terms: Avalanche Photodiode, Silicon Photomultiplier, Detector, Sensor

Abbreviation or Symbol: PMT

Synonym:

Reference(s): https://doi.org/10.1016/B0-12-369397-7/00529-X,

https://hamamatsu.magnet.fsu.edu/articles/photomultipliers.html

Alternative definition:

Related terms to alternative definition:

Reference(s) for alternative:

Originally drafted by: Nick Ross

Reviewed (or Edited) by:  D Busch, Heidrun Wabnitz, Edgar Guevara, Stefan Carp

Status: Open for review

#### Power

Definition: Radiant power refers to the amount of energy per unit time emitted by electromagnetic radiation, measured in watts (W) or decibels relative to a milliwatt (dBm)

Category: Hardware

Related terms: source, detector

Abbreviation or Symbol:

Synonym:

Reference(s): https://doi.org/10.1089/pho.2009.2503

Alternative definition:

Related terms to alternative definition:

Reference(s) for alternative:

Originally drafted by: M. Atif Yaqub

Reviewed (or Edited) by:  Edgar Guevara, Nick Ross,  Heidrun Wabnitz, Stefan Carp, Samuel Montero-Hernandez

Status: Complete (agreement from at least 5 team members)

#### Power supply

Definition: Power supply is an electronic device that provides a stable voltage, current, and frequency required for other devices to function correctly. This device maintains a stable rated wattage output consistently.

Category: Hardware

Related terms:

Abbreviation or Symbol:

Synonym:

Reference(s): https://doi.org/10.1109/TPEL.2019.2941528

Alternative definition:

Related terms to alternative definition:

Reference(s) for alternative:

Originally drafted by: M. Atif Yaqub, Nick Ross

Reviewed (or Edited) by:  Talukdar Raian Ferdous, Nick Ross, Usama A Syed, Samuel Montero-Hernandez, Stefan Carp

Status: Complete (agreement from at least 5 team members)

#### Probe

Definition: A probe or probeset is the arrangement of optodes on the head. Probes can be placed according to the international 10-20 system and/or other anatomical landmarks. The optodes (emitters and detectors) are typically located in a grid with fixed inter-optode-distances for long-distance and short-distance channels.

Category: Hardware

Related terms: topographic layout, optode placement, fNIRS caps, inter-optode-distance

Abbreviation or Symbol:

Synonym: probeset

Reference(s): -

Alternative definition:

'Probe' is widely used as a synonym for 'optode', e.g. in the standard IEC 80601-2-71:2015. Though technically not wrong, it might cause misunderstanding depending on the target group, as a 'probe' in the medical field refers to a sensor that is put INTO an object/test specimen (also see Oxford / Cambridge dictionary). ‘Optode’ should be preferred.

Related terms to alternative definition: Optode, Sensor, Detector, Source

Reference(s) for alternative:  https://dictionary.cambridge.org/dictionary/english-german/probe

Originally drafted by: Christina Artemenko

Reviewed (or Edited) by:  Franziska Klein, Christian Bartkowski, Edgar Guevara. Heidrun Wabnitz, Nick Ross

Status: Complete (agreement from at least 5 team members)

## Q

## R

#### Response Time

Definition: Time required for the step response of the equipment to move from its specified percentage of the final steady-state value to the other specified percentage. Response time is conventionally denoted by the rise time or fall time that represents the interval between the times corresponding to 10 % and 90 % of the step response amplitude during the transition.

Category: Hardware

Related terms:

Abbreviation or Symbol:

Synonym:

Reference(s): IEC 80601-2-71:2015, 201.3.215

Alternative definition:

Related terms to alternative definition:

Reference(s) for alternative:

Originally drafted by: IEC 80601-2-71

Reviewed (or Edited) by:  Heidrun Wabnitz

Status: No review (ISO/IEC term)

## S

#### Sampling frequency

Definition: The number of samples per second collected from a continuous signal through analog-to-digital conversion.

Category: Hardware

Related terms:

Abbreviation or Symbol:

Synonym: Sampling Rate

Reference(s): https://doi.org/10.3389/fpsyg.2017.00093

Alternative definition: Number of data samples collected per unit of time (in Hz) from each channel.

Related terms to alternative definition:

Reference(s) for alternative: https://doi.org/10.1117/1.NPh.8.1.012101

Originally drafted by:  Edgar Guevara; OntoNIRS (alternative definition)

Reviewed (or Edited) by:  Franziska Klein, Aykut Eken, Christian Bartkowski, Ann-Christine Ehlis, Thomas Dresler, Christina Artemenko, Vanessa Reindl, Murat Can Mutlu, Simone Cutini, Jason Longhurst, Haiyan Kong

Status: Complete (agreement from at least 5 team members)

#### Saturation

Definition: In terms of a photodiode, if the incident optical power continues to increase but the photocurrent output stops growing, it indicates that the photodiode has reached saturation. The saturation limit can be adjusted by reducing the resistance seen by the load or changing the reverse bias voltage applied. The saturation point represents the upper boundary of the  photodiode’s linear range, while the lower boundary of that linear range is determined by the noise. Beyond this range, the photodiode operates in a nonlinear manner.

Category: Hardware

Related terms: Photodiode

Abbreviation or Symbol: Psat

Synonym: photocurrent output limit

Reference(s): https://doi.org/10.1364/OE.422478

https://www.thorlabs.com/images/TabImages/Photodetector_Lab.pdf

Alternative definition:

Related terms to alternative definition:

Reference(s) for alternative:

Originally drafted by: Nick Ross

Reviewed (or Edited) by:  Michael Lührs, Franziska Klein, Edgar Guevara, Samuel Montero-Hernandez, Stefan Carp

Status: Complete (agreement from at least 5 team members)

#### Sensor

Definition: In the context of fNIRS, a sensor typically refers to a device, such as an avalanche photodiode or photomultiplier tube, that detects the near-infrared light that has passed through or reflected from the measured tissue. Some fNIRS systems additionally incorporate other physiological sensors, such as electroencephalography (EEG) or electrocardiography (ECG), and motion sensors like accelerometers or gyroscopes. These sensors capture other physiological signals that may be relevant to the study, providing complementary information to fNIRS measurements. By offering a more comprehensive understanding of brain activity and associated physiological responses, these sensors can contribute to improving accuracy and interpretation of the results.

Category: Hardware

Related terms: Optode, Detector, Motion Sensor

Abbreviation or Symbol:

Synonym:

Reference(s): National Cancer Institute. National Cancer Institute Thesaurus, 2017. Documentation page available on http://nciterms.nci.nih.gov.

https://doi.org/10.3390/photonics6030087

Alternative definition:

Related terms to alternative definition:

Reference(s) for alternative:

Originally drafted by: Christian Bartkowski

Reviewed (or Edited) by:  Christian Bartkowski, Nick Ross, Edgar Guevara, Mohinish Shukla, Franziska Klein

Status: Complete (agreement from at least 5 team members)

#### Short-separation channel

Definition:  A short separation channel is generated by placing a source and a detector at a short distance, in case of the adult brain, this is typically ~ 1 cm or less, in contrast to regular long-separation channels with ~3 cm source-detector distances. At such distances, channels are mostly sensitive to systemic activity resulting from extracerebral tissues (i.e., scalp and skull). The signal from these channels can be used to regress out extra-cerebral contaminations from the long-separation fNIRS signal in order to isolate the cortical functional response.

Category: Hardware

Related terms: Extra-cerebral signal, Source-detector separation, Measurement Channel

Abbreviation or Symbol: SSC, SDC

Synonym: short-distance channel, short channel

Reference(s): e.g. https://doi.org/10.1016/j.neuroimage.2013.04.070
https://doi.org/10.1117/1.NPh.2.2.025005

https://doi.org/10.1117%2F1.NPh.3.3.031405

Alternative definition:

Related terms to alternative definition: multi-distance measurements

Reference(s) for alternative:

Originally drafted by: Jörn Martin Horschig, Oliver Seidel-Marzi

Reviewed (or Edited) by:  Ann-Christine Ehlis, Franziska Klein, Vanessa Reindl, Jason Longhurst, Nick Ross, Christian Bartkowski, Edgar Guevara, Nick Ross, D Busch, M Atif Yaqub, Turgut Durduran, Alessandro Torricelli

Status: Complete (agreement from at least 5 team members)

#### Silicon Photodiode

Definition: A silicon photodiode (SiPD) is a silicon solid-state device that takes photons in and puts electrons out, usually at the rate of about 5 electrons out for every 10 photons. It produces current when photons impinge upon it; a silicon photodiode does not require an external power source in order to operate. SiPDs are commonly used as detectors in fNIRS systems because they are sensitive to light in the near-infrared spectrum. They are compact, fast, and have low noise, making them well-suited for use in wearable fNIRS systems besides laboratory systems.

Category: Hardware

Related terms: Detector, Sensor

Abbreviation or Symbol: SiPD

Synonym:

Reference(s): https://doi.org/10.1364/AO.39.006466

Alternative definition:

Related terms to alternative definition:

Reference(s) for alternative:

Originally drafted by: OntoNIRS, Christian Bartkowski

Reviewed (or Edited) by:  Franziska Klein, Christian Bartkowski, D Busch, Nick Ross M. Atif Yaqub, Samuel Montero-Hernandez

Status: Complete (agreement from at least 5 team members)

#### Silicon Photomultiplier

Definition: The silicon photomultiplier (SiPM) is a type of sensor/detector with compactness, wide dynamic range, fast timing, and high gain factors. SiPMs are solid-state photodetectors made of an array of hundreds or thousands of integrated single-photon avalanche diodes (SPADs), called microcells or pixels whose outputs are merged into one, and therefore result in high sensitivity, wide dynamic range, and gain factors of 105 to 106, while allowing small construction sizes to even 1.1 x 1.1 mm2 frame sizes.

Upon detecting a photon, the SPAD generates a large electrical output signal due to internal avalanche multiplication. The SiPM allows one to detect and count photons with high resolution and single photon sensitivity. The internal avalanche amplification is also fast enough to obtain excellent timing information on the arrival time of the detected photons within several tens of picoseconds.

However, the supply voltage depends on the APD technology used. It typically varies between 20 V and 100 V, thus being from 15 to 75 times lower than the voltage required for a photomultiplier tube's (PMT) operation but also 5 to 20 times higher than for a SiPD. These high voltages require additional safety measures for applications on the human head. Furthermore, like APDs, SiPMs have a high temperature sensitivity and increased noise components. However, SiPMS have a smaller dynamic range and saturation occurs at lower light levels when compared to APDs

Category: Hardware

Related terms: avalanche photodiode, APD, silicone photodiode, SiPD, detector, sensor, single photon avalanche diodes, SPAD, PMT, photomultiplier tube

Abbreviation or Symbol: SiPM

Synonym: solid-state photomultiplier (SSPM), or multi-pixel photon counter (MPPC)

Reference(s): https://doi.org/10.1088/1361-6560/ab7b2d

https://doi.org/10.1142/9789812776464_0101

Alternative definition:

Related terms to alternative definition:

Reference(s) for alternative:

Originally drafted by: Christian Bartkowski

Reviewed (or Edited) by:  Franziska Klein, Nick Ross, Alessandro Torricelli, Stefan Carp

Status: Open for review

#### Smart Textiles

Definition: Smart textiles (ST) are fabrics-based products that can sense (Passive ST), react (Active ST), and even adapt (Ultra ST) to their environment, whereby stimuli can have an electrical, photonic, thermal, chemical, magnetic, etc. origin. For this, they are fitted with electronic components to process information. The decision can occur locally in embedded electronic devices or remotely via connected external systems.

Category: Hardware

Related terms: headgear, headband, caps, patches, armband, shirt, detectors, sources

Abbreviation or Symbol: ST

Synonym: Smart Clothing, Smart Garments, Smart Fabrics

Reference(s): https://doi.org/10.1007/978-3-319-50124-6_1

https://shop.elsevier.com/books/smart-textiles-and-their-applications/koncar/978-0-08-100574-3

https://doi.org/10.3390/s140711957

Alternative definition:

Related terms to alternative definition:

Reference(s) for alternative:

Originally drafted by: Christian Bartkowski

Reviewed (or Edited) by:  M. Atif Yaqub, Samuel Montero-Hernandez, Nick Ross

Status: Open for review

#### Source

Definition: Source is an optoelectronic element that delivers light (e.g. a laser or an LED). In fNIRS, it refers to a source of near-infrared light that is used to illuminate a tissue being interrogated. Sources can either require optic fibers to carry the light to source optodes placed on the scalp or can be integrated into optodes, wearable probes or modules placed directly on the scalp.

Category: Hardware

Related terms: detector, source-detector pair, channel, probe, diode

Abbreviation or Symbol: S, Tx, SRC

Synonym: emitter

Reference(s): https://doi.org/10.1146/annurev-psych-010213-115108

https://doi.org/10.1016/S0006-3223(02)01550-0

https://doi.org/10.3390/s21186106

https://www.intechopen.com/chapters/28786

Alternative definition:

Related terms to alternative definition:

Reference(s) for alternative:

Originally drafted by: Mohinish Shukla, Ola Dopierala, Heidrun Wabnitz, Meryem Ayşe Yücel

Reviewed (or Edited) by:  Talukdar Raian Ferdous, Edgar Guevara, Christian Bartkowski, Nick Ross, Franziska Klein, Murat Can Mutlu, Heidrun Wabnitz, Alessandro Torricelli

Status: Complete (agreement from at least 5 team members)

#### Source-detector pair

Definition: A set of two optodes (one light source and one detector) which is considered as a single measurement unit (i.e., measurement channel). Although in theory any source can be paired with any detector, most systems/researchers set up a region of interest- and a distance-criteria (how far away from each other are the two optodes located, i.e., source-detector separation) to define source-detector pairs that interrogate the brain. A single optode (either source or detector) can be part of multiple source-detector pairs.

Category: Hardware

Related terms: source, detector, array, optode, montage, High Density (Source-detector Array), measurement channel

Abbreviation or Symbol: SD, S-D

Synonym: SD pair, measurement channel, S-D combination

Reference(s): https://doi.org/10.1118/1.1354627

https://doi.org/10.1038/s41598-018-21716-z

Alternative definition:

Related terms to alternative definition:

Reference(s) for alternative:

Originally drafted by: Ola Dopierala

Reviewed (or Edited) by:  Edgar Guevara, Christian Bartkowski, Nick Ross, Heidrun Wabnitz, Samuel Montero-Hernandez

Status: Complete (agreement from at least 5 team members)

#### Source-detector separation

Definition: Source-detector separation (SDS) refers to the physical distance between each fNIRS source-detector pair of optodes. It determines the signal-to-noise ratio (SNR), the number of measurable photons, depth sensitivity to the change in hemoglobin, the allowed amount of sources/detectors/channels, and the size of covered regions. Depending on the population being studied (e.g., adults vs. infants) and the purpose of the channels (e.g., measuring oxygenated hemoglobin in the brain vs. removing the confounding effects of superficial tissue physiology), the optimal SDS will vary.  In a reflection geometry, larger SDS increases depth sensitivity (i.e., to deep regions further from the surface), at the cost of reduced total detected intensity.

Category: Hardware

Related terms: short-channel separation, source-detector pair

Abbreviation or Symbol: SDS

Synonym: Source-detector separation range, Source-detector separation, inter-optode distance

Reference(s): https://doi.org/10.1117/1.NPh.2.2.025005

https://doi.org/10.1111/cpf.12156

https://doi.org/10.1117/1.JBO.21.9.091307

Alternative definition:  The inter-optode distance is the spatial distance between a source and a detector. In circular geometries it is defined as the length of the chord between the optodes.

Related terms to alternative definition:

Reference(s) for alternative:  https://doi.org/10.1007/978-1-4615-3404-4_17

Originally drafted by: Chia-Chuan Yu

Reviewed (or Edited) by:  Franziska Klein, Aykut Eken, Christian Bartkowski, Talukdar Raian Ferdous, D Busch, Nick Ross, Edgar Guevara, Alessandro Torricelli, Heidrun Wabnitz

Status: Complete (agreement from at least 5 team members)

#### Source Optode

Definition: Source or injection optode is a common term which often identifies a part of an fNIRS system that is in contact with the surface of the tissue being measured. A source optode either directly contains the light sources (e.g. for an LED-based system), or can hold the distal end of source fibers (e.g. for a laser-based system) emitting light at two or more wavelengths. This component may also contain a light guide, a lens, and a housing.

Category: Hardware

Related terms: Source, Detector, emitter, optode

Abbreviation or Symbol: S, Tx, SRC

Synonym: Emitter Optode

Reference(s):

Alternative definition:

Related terms to alternative definition:

Reference(s) for alternative:

Originally drafted by: Heidrun Wabnitz, Meryem Ayşe Yücel

Reviewed (or Edited) by:  Edgar Guevara, Christian Bartkowski, Heidrun Wabnitz, Nick Ross, Alessandro Torricelli

Status: Complete (agreement from at least 5 team members)

## T

#### Time multiplexing

Definition: Time multiplexing is a method of transmitting and receiving two or more independent signals over a common channel path by means of synchronized switches at each end of the transmission line. This way each signal appears on the line only a fraction of time in an alternating pattern. In fNIRS this technique of switching on and off different light sources at fixed intervals avoids the superposition of two sources being detected by a single detector simultaneously. This strategy is commonly employed in fNIRS devices to enable the detector to identify from which emitter the light came from. The temporal resolution of a device is inversely proportional to the number of time multiplexing steps and step width (i.e., the more steps there are, the slower the frequency of acquisition).

Category: Hardware

Related terms: Frequency multiplexing

Abbreviation or Symbol:

Synonym: Time division multiplexing

Reference(s): DOI: 10.1117/1.nph.8.1.012101; DOI: 10.1201/b17289 (section 7.3.1)

Alternative definition:

Related terms to alternative definition:

Reference(s) for alternative:

Originally drafted by: Rickson Mesquita, Anna Exner

Reviewed (or Edited) by:  Rickson Mesquita, Anna Exner, Guy Perkins, Stefan Carp, Samuel Montero-Hernandez

Status: Complete (agreement from at least 5 team members)

#### Transmitter

Definition: In wearable/portable fNIRS systems, a transmitter refers to the communication unit worn by the subject. It may include sensors that record fNIRS data and other physiological data such as heart rate or motion. The recorded data can then be transferred wirelessly to a receiver unit using a variety of communication protocols, such as Bluetooth or Wi-Fi.

Category: Hardware

Related terms:

Abbreviation or Symbol: Tx

Synonym:

Reference(s): https://doi.org/10.1109/EMBC.2013.6609962

https://doi.org/10.3390/brainsci10020085

https://doi.org/10.1002/cta.3047

https://publications.polymtl.ca/1346/1/2014_EhsanKamrani.pdf

Alternative definition:

Related terms to alternative definition:

Reference(s) for alternative:

Originally drafted by: Jörn Martin Horschig

Reviewed (or Edited) by:  Edgar Guevara, Christian Bartkowski, D Busch, Nick Ross, Samuel Montero-Hernandez

Status: Complete (agreement from at least 5 team members)

## U

## V

#### Vertical Cavity Surface Emitting Laser

Definition: A Vertical Cavity Surface Emitting Laser (VCSEL) is a light-emitting, laser-based component that provides high wavelength stability with a very stable, short cavity. VCSEL of 3 or 4” wafers have less than a 2 nm standard deviation for cavity wavelength. A VCSEL operates in a single longitudinal mode, which avoids signal distortions due to mode beatings. It has low power consumption, and the circular output beam allows simpler optics than side-emitting lasers. Its temperature sensitivity of wavelengths is ~5 times less sensitive to temperature variations than edge emitters. Over a 20°C change in temperature causes the emission wavelength to vary less than 1,4 nm (edge emitters = ~7 nm).

A VCSEL allows highly compact designs; it enables the packaging of multi-wavelength source chips with reduced placing shifts. However, as VCSEL’s typical use field is in data transfer and heating systems, multi-wavelength emitters between 600 and 900 nm are not widely available to the public.

Category: Hardware

Related terms: source, multi-wavelength, emitter, laser

Abbreviation or Symbol: VCSEL

Synonym:

Reference(s): https://doi.org/10.1117/12.2209320

https://doi.org/10.1117/12.2077209

https://doi.org/10.1117/12.2210909

https://doi.org/10.1117/1.3462996

Alternative definition:

Related terms to alternative definition:

Reference(s) for alternative:

Originally drafted by: Christian Bartkowski

Reviewed (or Edited) by:  Franziska Klein, Nick Ross

Status: Open for review

## W

## X

## Y

## Z